## Item 1(4)
Lot 1: Measuring Equipment
Item 1(4): Taurus TFR Measuring Instrument
Guaranteed Technical Particular (GTP) to be filled by bidders.
SL# | Description | Technical Specifications | Bidders to fill in (Comply/Not)
[Insert Bidder's name here]
0 | Make, Model and Country of origin
1 | Functional Requirement | The equipment should be able to measure the impedance, resistance and inductance of the tower foot grounding without isolating the top earth wire, in online condition. Instrument shall be immune to noise and induction
2 | Special Features | Portable, light weight, inbuilt rechargeable battery operated Automatic suppression of interference.
3 | Voltage Output | 250 Volts
4 | Current output | The equipment should be able to give constant output of 20mA upto a load of 2 kohms with high output open circuit voltage of 250 volts
5 | Test Frequency | 25 KHz as per IEEE
6 | Measurement Range | 0.05 ohms to 2 k Ohms
7 | Short Circuit Current | 20 mA Automatic
8 | Measurement Accuracy | +/-2 % to 5%
9 | Display | Graphic LCD Display
10 | Power Supply | 15V DC adapter with AC 230V Single phase AC, with in-built Rechargeable battery. Battery capacity of 12V,7 Ah, capable of working continuously for min 6 hrs
11 | Repeatability | The instrument should be proven for repeatability of reading in charged upto 765kV AC and 800Kv DC lines
12 | Memory | Instrument should have internal memory storage up to 2000 results
13 | GPS Location | Should locate Lat-Long of a tower without using
Tele communication sim card. Should be supplied
along with Laptop / tab.
14 | Mobile Software | Instrument should be supplied with advanced Android application to store the data and GPS integration.
Software should have facility to enter tower number & store the data, Latitude & longitude locations automatically in real time.
15 | Calibrator | Kit should be supplied along with standard calibrator to verify the readings of Impedance, Resistance & Inductance.
Calibrator should be certified by NABL Labs.
16 | Interface | RS 232/Bluetooth
17 | Peg Check Indicator | Should be provided indication/alarm for error free operation
18 | Accessories | a. Galvanized spikes – 3 no’s,
b. 5 m,25m,50m test leads- 1 set
c. 2 m,9m,15m test leads- 1 set
d. AC adapter, Hammer and soft carrying case, manual which are required for testing.
19 | Environment | Temp- -5 d.c to 50 d.c.
Storage Temp: -20 d.c to 50 d.c
Humidity – 95%RH non Condensing
20 | Calibration certificate | Calibration certificate to be provided from any NABL accredited laboratory
21 | Warranty | The warranty period shall be 12 months for any manufacturing defects.
*The bidders are required to submit GTP in both PDF and Excel formats.

## Item 2(3)
Lot 2: Electrical Equipment
Item 2(3): Transformer Oil Filtration Plant Complete Set
Guaranteed Technical Particular (GTP) to be filled by bidders.
Parameters |  | Offer
[Insert Bidder's name here]
Model No.
Manufacturer/ Brand
Capacity(L/min)
Vacuum pressure
Working pressure
Ultimate vacuum
Temperature range
Breakdown voltage
Parameter after purification | Water content
Gas content
Filtering rate
Filtering precision
Continuous work
No failure running
Power supply
Working noise
Heating power (kw)
Total power (kw)
*The bidders are required to submit GTP in both PDF and Excel formats.

## Item 2(18)
Lot No. 2: Measuring Equipment
Item 2(18): Meter Testing Equipment
Guaranteed Technical Particular (GTP) to be filled by bidders.
SL# | Item | Parameters | Bidders to specify
[Insert bidder's name here]
1 | Makers name and country | Please specify manufacturer’s name
2 | Type of meter
3 | Accuracy Class | 3.1 Direct mode:
3.2 Clamp-on mode:
Error compensated CT clamps with compensation as part of CT
4 | Parameters displayed (YES/NO) | 4.1 Instantaneous voltage, U of every phase
4.2 Instantaneous line current, U of each phase
4.3 Simultaneous display of active (P), reactive (Q) and apparent power (S)
4.4 Instantaneous power factor of each phase & total power factor.
4.5 Instantaneous frequency
4.6 Phase sequence
4.7 Active, reactive (lag/lead) and apparent three phase energies
4.8 Continuous update of active, reactive (lead/lag) and apparent energies on display.
4.9 Time
4.10 Vectorial graphic display
4.11 Harmonic analysis display of U, I, P, Q
4.12 Real time waveform display of selected voltage & current.
5 | Display Resolution (SPECIFIC VALUES) | 5.1 Voltage:
5.2 Current:
5.3 Power factor:
5.4 Energy:
5.5 % Error Resolution:
6 | Connection check (YES/NO) | 6.1 Missing potential
6.2 Missing current
6.3 Reverse current if any current is reverse
6.4 Phase sequence “Forward or Reverse”
6.5 Over Current
6.6 Over Voltage
7 | Type of Display (DESCRIBE) | Specify the type of display and the resolution.
8 | Interface (YES/NO) | 8.1 USB or RJ45 connector for connecting to the PC
8.2 Scanning head
8.3 Remote snap switch to count pulses
8.4 Wi-Fi Connectivity
8.5 Webserver for remote connectivity
9 | Memory (YES/NO) | 9.1 Capacity of minimum 20000 test results
9.2 Upload of customer database for at least 5 customers
9.3 Memory management, part of instrument & external software
9.4 Selection of special attributive tests
10 | Instantaneous parameters to be logged in memory during each test | 10.1 Customer information
10.2 Site information
10.3 Attributive Test Information
10.4 Measurement results
11 | Auto detection of current measurement input (YES/NO) | Automatic detection of DIRECT, 120A CT modes
12 | Functional Checks (DESCRIBE) | 12.1 Accuracy using scanning head (Describe)
12.2 Register Test
12.3 Data storage function.
13 | Scanning head (YES/NO) | Common for rotor mark & LED pulses to sense pulses up to 500Hz.
14 | Snap switch | Snap switch to operate equipment remotely.
15 | Dial test facility | Test of Active, Reactive and Apparent Energy, simultaneously
Test of Dial/Counter of Meter under test for all 3 vectors (Active, Reactive and Apparent) in one shot or simultaneously
16 | Keyboard (YES/NO) | Alphanumeric type
17 | Carrying case | Describe the type / material of the carrying case.
18 | Weight in kg | Without case:
With case:
19 | Type Test | Calibration report from recognized lab shall be submitted.
Declaration from manufacturer for standards adopted for test of insulation influence quantities (CE Declaration).
*The bidders are required to submit GTP in both PDF and Excel formats.

## Lot 6(1)
Lot 6: CTPT Metering Units
Guaranteed Technical Particular (GTP) to be filled by bidders for (1) For 11kV & 33kV single phase Current Transformers
S.No. | PARTICULARS |  | TO BE FURNISHED BY BIDDER
[Insert Bidder's name here]
1 | NAME & ADDRESS OF MANUFACTURER.
2 | MANUFACTURER'S TYPE & DESIGN
3 | NOMINAL SYSTEM VOLTAGE (KV rms)
4 | HIGHEST SYSTEM VOLTAGE (KV rms)
5 | INSULATION LEVEL :
a) IMPULSE WITHSTAND VOLTAGE
(KV PEAK)
b) ONE MINUITE POWER
FREQUENCY DRY WITHSTAND VOLTAGE (KVrms)
(i) PRIMARY(KVrms)
(ii) SECONDARY (KVrms)
6 | RATED FREQUENCY (Hz)
7 | TRANSFORMATION RATIO.
8 | RATED OUTPUT (VA BURDEN)
9 | CLASS OF ACCURACY.
10 | DETAIL OF WINDINGS.
| PARTICULARS | CT RATIO | NO.OF TURNS | CROSS SECTIONAL AREA OF EACH TURN IN SQ.MM. | TOTAL WEIGHT OF WDG. IN Kg. PER
|  |  |  | (Approx.) | UNIT
| i) PRIMARY | 15-50/1 A
25-75/1 A
| ii) SECONDARY | 15-50/1 A
25-75/1 A
11 | GUARANTEED RATIO ERROR (MAX.)
GUARANTEED PHASE ANGLE ERROR (MAX.)
*The bidders are required to submit GTP in both PDF and Excel formats.

## Lot 6(2)
Lot 6: CTPT Metering Units
Guaranteed Technical Particular (GTP) to be filled by bidders for (2) 11kV & 33kV three phase potential transformers
S.No. | PARTICULARS | TO BE FURNISHED BY BIDDER
[Insert bidder's name here]
1 | NAME & ADDRESS OF
MANUFACTURER.
2 | MANUFACTURER'S TYPE & DESIGN.
3 | NOMINAL SYSTEM VOLTAGE (KV rms)
4 | HIGHEST SYSTEM VOLTAGE (KV rms)
5 | INSULATION LEVEL :
a) IMPULSE WITHSTAND
VOLTAGE(KV PEAK)
b) ONE MINUITE POWER FREQUENCY WITHSTAND VOLTAGE (KV PEAK)
6 | RATED FREQUENCY (Hz)
7 | RATED TRANSFORMATION RATIO.
8 | RATED OUTPUT (VA BURDEN PER PHASE)
9 | CLASS OF ACCURACY.
10 | RATED VOLTAGE FACTOR AND TIME.
11 | WINDING CONNECTIONS
a) PRIMARY
b) SECONDARY
12 | DETAIL OF WINDINGS.
| Particulars | PT Ratio | No. of Turns | Cross-Sectional | Total Wt. of
|  |  |  | Area of Each (Sq.mm) | winding in KG (Approx.)
i) PRIMARY WINDING (COPPER).
ii) SECONDARY WINDING (COPPER). 110 V
13 | GURANTEED RATIO ERROR (MAX.)
14 | GURANTEED PHASE ANGLE ERROR (MAX.)
15 | GURANTEED MAX. TEMP. RISE OF THE WINDING OVER AN AMBIENT TEMP. OF 50 DEG. C AT RATED CONTINOUS THERMAL CURRENT, AT RATED FREQUENCY & WITH RATED BURDEN.
*The bidders are required to submit GTP in both PDF and Excel formats.